## Dashboard
COMPONENT QC INSPECTION — QUALITY DASHBOARD
Lab Sensor & Sistem Telekontrol (SSTK) ¦ Dept. Teknik Nuklir & Teknik Fisika, UGM ¦ 4 Praktikum ¦ ~80 Mahasiswa/Minggu ¦ All values auto-calculated
| Pass Rate Elektronika Analog |  | Pass Rate Elektronika Digital |  |  |  | Pass Rate Sistem Instrumentasi | Pass Rate Jaringan Komunikasi
| 87.5% |  | 70.83% |  |  |  | 81.25% | 62.5%
▶ CROSS-PRACTICUM INSPECTION SUMMARY
| Practikum | Total Inspected | Total PASS |  |  | Top Defect Count
| Elektronika Analog | 48 | 42 | 6 | 87.5% | 6
| Elektronika Digital | 48 | 34 | 14 | 70.83% | 14
| Sistem Instrumentasi | 48 | 39 | 9 | 81.25% | 9
| Jaringan Komunikasi | 48 | 30 | 18 | 62.5% | 18
| TOTAL ALL PRACTIKUM | 192 | 145 | 47 | 75.52%
📋 FORMULA & CALCULATION NOTES — DASHBOARD FORMULA NOTES
Column / Field | Calculation Logic
KPI Cards — Pass Rate | [Formula] (relay_pass + switch_pass) / (relay_total + switch_total) |  |  |  |  | AUTO. Pulls PASS and total counts from each practikum sheet's summary row. Updates automatically when new inspection data is added.
Summary Table — Pass Rate | [Formula] pass_count / total_inspected per practikum |  |  |  |  | AUTO. Per-practikum pass rate from each inspection log summary. All 4 values feed into the cross-practikum bar chart.
Summary Table — Total/PASS/FAIL | [Formula] direct reference to practikum sheet summary row |  |  |  |  | AUTO. Values sourced directly from summary formulas in each inspection sheet. No manual input required on this sheet.
Bar Chart — Pass Rate by Practikum | References Summary Table Col F (Pass Rate) |  |  |  |  | AUTO-updated chart. Bars reflect current pass rates from all 4 practikum inspection logs. Provides at-a-glance comparison of quality across practikum types.
ℹ Dashboard requires zero manual input. All values auto-calculated from the 4 practikum inspection sheets. Refresh Excel or open file to recalculate.

## Elek_Analog
PRAKTIKUM ELEKTRONIKA ANALOG — COMPONENT QC INSPECTION LOG
Lab Sensor & Sistem Telekontrol (SSTK) ¦ Dept. Teknik Nuklir & Teknik Fisika, UGM ¦ ~20 Mahasiswa/Sesi
Test Equipment: Multimeter (Sanwa CD771), LCR Meter, Oscilloscope (GDS-1022), DC Power Supply
No. | Part Number | Component Description | Batch / Lot No. | Inspection Date | Qty | Visual Inspection | Spec / Rated Value | Measured Value | Functional Test 1 | Test 1 Result | Functional Test 2 | Test 2 Result | Overall Result | Remarks / Disposition
⚠ Columns shaded yellow = manual input per inspection session ¦ Columns shaded gray = auto-calculated by formula ¦ Overall Result = auto (PASS only if ALL tests pass)
Inspection Session: Sesi A — Week 1 ¦ Batch: LOT-A-2024
1 | R-10kΩ-CFR | Resistor 10kΩ | LOT-A-581 | 05-Apr-2026 | 24 | Pass | 10 | 10.19 | Measured Value | Pass | Tolerance | Pass | PASS | Release to practicum
2 | R-1kΩ-CFR | Resistor 1kΩ | LOT-A-875 | 05-Apr-2026 | 25 | Pass | 1 | 0.94 | Measured Value | Pass | Tolerance | Pass | PASS | Release to practicum
3 | C-100µF-EC | Capacitor 100µF | LOT-A-462 | 05-Apr-2026 | 22 | Pass | 100 | 105.47 | Measured Cap | Pass | ESR Test | Pass | PASS | Release to practicum
4 | C-10nF-MLC | Capacitor 10nF | LOT-A-656 | 05-Apr-2026 | 21 | Pass | 10 | 10.05 | Measured Cap | Pass | Tolerance | Pass | PASS | Release to practicum
5 | D-1N4007 | Rectifier Diode | LOT-A-805 | 06-Apr-2026 | 20 | Pass | 0.7 | 0.69 | Vf Measured | Pass | Continuity | Pass | PASS | Release to practicum
6 | D-1N4148 | Signal Diode | LOT-A-644 | 06-Apr-2026 | 23 | Pass | 0.6 | 0.59 | Vf Measured | Pass | Continuity | Fail | FAIL | Quarantine — do not issue to students
7 | Q-BC547 | NPN Transistor | LOT-A-743 | 06-Apr-2026 | 23 | Pass | 110 | 100.08 | hFE Measured | Pass | Vce_sat | Pass | PASS | Release to practicum
8 | Q-BC557 | PNP Transistor | LOT-A-499 | 06-Apr-2026 | 24 | Pass | 110 | 117.75 | hFE Measured | Pass | Vce_sat | Pass | PASS | Release to practicum
9 | U-LM741 | Op-Amp IC | LOT-A-424 | 07-Apr-2026 | 21 | Pass | 5 | 5.49 | VOS Measured | Pass | Pin Continuity | Pass | PASS | Release to practicum
10 | U-LM324 | Quad Op-Amp IC | LOT-A-141 | 07-Apr-2026 | 22 | Pass | 5 | 5.31 | VOS Measured | Pass | Pin Continuity | Pass | PASS | Release to practicum
11 | U-LM7805 | Voltage Reg +5V | LOT-A-964 | 07-Apr-2026 | 24 | Pass | 5 | 5.33 | Vout Measured | Pass | Load Reg Test | Pass | PASS | Release to practicum
12 | POT-10kΩ | Potentiometer 10k | LOT-A-728 | 07-Apr-2026 | 21 | Pass | 10 | 10.84 | Measured R | Pass | Sweep Test | Pass | PASS | Release to practicum
Inspection Session: Sesi B — Week 2 ¦ Batch: LOT-B-2025
1 | R-10kΩ-CFR | Resistor 10kΩ | LOT-B-341 | 12-Apr-2026 | 25 | Pass | 10 | 9.68 | Measured Value | Pass | Tolerance | Pass | PASS | Release to practicum
2 | R-1kΩ-CFR | Resistor 1kΩ | LOT-B-709 | 12-Apr-2026 | 23 | Pass | 1 | 0.91 | Measured Value | Pass | Tolerance | Pass | PASS | Release to practicum
3 | C-100µF-EC | Capacitor 100µF | LOT-B-443 | 12-Apr-2026 | 20 | Pass | 100 | 111.98 | Measured Cap | Fail | ESR Test | Pass | FAIL | Quarantine — do not issue to students
4 | C-10nF-MLC | Capacitor 10nF | LOT-B-800 | 12-Apr-2026 | 21 | Pass | 10 | 10.87 | Measured Cap | Pass | Tolerance | Pass | PASS | Release to practicum
5 | D-1N4007 | Rectifier Diode | LOT-B-694 | 13-Apr-2026 | 18 | Pass | 0.7 | 0.63 | Vf Measured | Pass | Continuity | Pass | PASS | Release to practicum
6 | D-1N4148 | Signal Diode | LOT-B-254 | 13-Apr-2026 | 21 | Pass | 0.6 | 0.56 | Vf Measured | Pass | Continuity | Pass | PASS | Release to practicum
7 | Q-BC547 | NPN Transistor | LOT-B-922 | 13-Apr-2026 | 25 | Pass | 110 | 115.15 | hFE Measured | Pass | Vce_sat | Pass | PASS | Release to practicum
8 | Q-BC557 | PNP Transistor | LOT-B-636 | 13-Apr-2026 | 19 | Pass | 110 | 114.33 | hFE Measured | Pass | Vce_sat | Pass | PASS | Release to practicum
9 | U-LM741 | Op-Amp IC | LOT-B-408 | 14-Apr-2026 | 22 | Pass | 5 | 4.56 | VOS Measured | Pass | Pin Continuity | Pass | PASS | Release to practicum
10 | U-LM324 | Quad Op-Amp IC | LOT-B-639 | 14-Apr-2026 | 24 | Pass | 5 | 4.66 | VOS Measured | Pass | Pin Continuity | Pass | PASS | Release to practicum
11 | U-LM7805 | Voltage Reg +5V | LOT-B-280 | 14-Apr-2026 | 18 | Pass | 5 | 4.67 | Vout Measured | Pass | Load Reg Test | Pass | PASS | Release to practicum
12 | POT-10kΩ | Potentiometer 10k | LOT-B-692 | 14-Apr-2026 | 19 | Pass | 10 | 10.43 | Measured R | Pass | Sweep Test | Pass | PASS | Release to practicum
Inspection Session: Sesi C — Week 3 ¦ Batch: LOT-C-2026
1 | R-10kΩ-CFR | Resistor 10kΩ | LOT-C-221 | 19-Apr-2026 | 21 | Pass | 10 | 10.04 | Measured Value | Pass | Tolerance | Pass | PASS | Release to practicum
2 | R-1kΩ-CFR | Resistor 1kΩ | LOT-C-750 | 19-Apr-2026 | 21 | Pass | 1 | 0.9 | Measured Value | Pass | Tolerance | Pass | PASS | Release to practicum
3 | C-100µF-EC | Capacitor 100µF | LOT-C-599 | 19-Apr-2026 | 20 | Pass | 100 | 105.56 | Measured Cap | Pass | ESR Test | Pass | PASS | Release to practicum
4 | C-10nF-MLC | Capacitor 10nF | LOT-C-732 | 19-Apr-2026 | 18 | Pass | 10 | 11.06 | Measured Cap | Fail | Tolerance | Pass | FAIL | Quarantine — do not issue to students
5 | D-1N4007 | Rectifier Diode | LOT-C-709 | 20-Apr-2026 | 18 | Pass | 0.7 | 0.65 | Vf Measured | Pass | Continuity | Pass | PASS | Release to practicum
6 | D-1N4148 | Signal Diode | LOT-C-225 | 20-Apr-2026 | 21 | Pass | 0.6 | 0.55 | Vf Measured | Pass | Continuity | Pass | PASS | Release to practicum
7 | Q-BC547 | NPN Transistor | LOT-C-693 | 20-Apr-2026 | 23 | Pass | 110 | 101.82 | hFE Measured | Pass | Vce_sat | Pass | PASS | Release to practicum
8 | Q-BC557 | PNP Transistor | LOT-C-379 | 20-Apr-2026 | 21 | Pass | 110 | 113.55 | hFE Measured | Pass | Vce_sat | Pass | PASS | Release to practicum
9 | U-LM741 | Op-Amp IC | LOT-C-509 | 21-Apr-2026 | 20 | Pass | 5 | 4.94 | VOS Measured | Pass | Pin Continuity | Pass | PASS | Release to practicum
10 | U-LM324 | Quad Op-Amp IC | LOT-C-314 | 21-Apr-2026 | 23 | Pass | 5 | 4.88 | VOS Measured | Pass | Pin Continuity | Pass | PASS | Release to practicum
11 | U-LM7805 | Voltage Reg +5V | LOT-C-566 | 21-Apr-2026 | 21 | Pass | 5 | 4.54 | Vout Measured | Pass | Load Reg Test | Pass | PASS | Release to practicum
12 | POT-10kΩ | Potentiometer 10k | LOT-C-870 | 21-Apr-2026 | 25 | Pass | 10 | 10.38 | Measured R | Pass | Sweep Test | Fail | FAIL | Quarantine — do not issue to students
Inspection Session: Sesi D — Week 4 ¦ Batch: LOT-D-2027
1 | R-10kΩ-CFR | Resistor 10kΩ | LOT-D-667 | 26-Apr-2026 | 18 | Pass | 10 | 10.73 | Measured Value | Pass | Tolerance | Pass | PASS | Release to practicum
2 | R-1kΩ-CFR | Resistor 1kΩ | LOT-D-622 | 26-Apr-2026 | 22 | Pass | 1 | 1.11 | Measured Value | Fail | Tolerance | Pass | FAIL | Quarantine — do not issue to students
3 | C-100µF-EC | Capacitor 100µF | LOT-D-782 | 26-Apr-2026 | 24 | Pass | 100 | 93.4 | Measured Cap | Pass | ESR Test | Pass | PASS | Release to practicum
4 | C-10nF-MLC | Capacitor 10nF | LOT-D-169 | 26-Apr-2026 | 23 | Pass | 10 | 9.12 | Measured Cap | Pass | Tolerance | Fail | FAIL | Quarantine — do not issue to students
5 | D-1N4007 | Rectifier Diode | LOT-D-414 | 27-Apr-2026 | 25 | Pass | 0.7 | 0.66 | Vf Measured | Pass | Continuity | Pass | PASS | Release to practicum
6 | D-1N4148 | Signal Diode | LOT-D-544 | 27-Apr-2026 | 19 | Pass | 0.6 | 0.62 | Vf Measured | Pass | Continuity | Pass | PASS | Release to practicum
7 | Q-BC547 | NPN Transistor | LOT-D-222 | 27-Apr-2026 | 19 | Pass | 110 | 113.61 | hFE Measured | Pass | Vce_sat | Pass | PASS | Release to practicum
8 | Q-BC557 | PNP Transistor | LOT-D-155 | 27-Apr-2026 | 22 | Pass | 110 | 120.39 | hFE Measured | Pass | Vce_sat | Pass | PASS | Release to practicum
9 | U-LM741 | Op-Amp IC | LOT-D-336 | 28-Apr-2026 | 21 | Pass | 5 | 5.05 | VOS Measured | Pass | Pin Continuity | Pass | PASS | Release to practicum
10 | U-LM324 | Quad Op-Amp IC | LOT-D-683 | 28-Apr-2026 | 22 | Pass | 5 | 4.92 | VOS Measured | Pass | Pin Continuity | Pass | PASS | Release to practicum
11 | U-LM7805 | Voltage Reg +5V | LOT-D-544 | 28-Apr-2026 | 19 | Pass | 5 | 5.32 | Vout Measured | Pass | Load Reg Test | Pass | PASS | Release to practicum
12 | POT-10kΩ | Potentiometer 10k | LOT-D-303 | 28-Apr-2026 | 20 | Pass | 10 | 10.18 | Measured R | Pass | Sweep Test | Pass | PASS | Release to practicum
SESSION SUMMARY
Metric | [5 empty] | Value | Total Records | PASS | FAIL | Pass Rate
[7 empty] | 48 | 42 | 6 | 87.5%
📋 FORMULA & CALCULATION NOTES — ELEK ANALOG
Column / Field | Calculation Logic
Col K — Test 1 Result | [Formula] IF(g="Fail","Fail", IF(ABS(i-h)/h ≤ 10%, "Pass", "Fail")),TRUE())),TRUE())),TRUE())),TRUE())) | [5 empty] | AUTO. Fails if Visual (G) already failed. Otherwise, passes only if measured value (I) is within ±10% of the rated spec (H). Tolerance 10% covers standard component variability.
Col N — Overall Result | [Formula] IF(AND(g="Pass", k="Pass", m="Pass"), "PASS", "FAIL") | [5 empty] | AUTO. Component passes only if ALL three tests pass: Visual (G), Test 1 (K), and Test 2 (M). Any single failure triggers FAIL — consistent with incoming inspection standards.
Col O — Remarks / Disposition | [Formula] IF(n="PASS", "Release to practicum", "Quarantine — do not issue") | [5 empty] | AUTO. Directly derived from Overall Result. PASS components are cleared for student use. FAIL components are quarantined and tracked in the Defect Log sheet.
Col G — Visual Inspection | [MANUAL INPUT: Pass / Fail] | [5 empty] | Inspector physically checks component body, pin integrity, marking, and housing. Recorded as Pass or Fail. Basis for all downstream formula calculations.
Col H — Spec / Rated Value | [MANUAL INPUT: numeric] | [5 empty] | Rated value from component datasheet or procurement spec (e.g., 10kΩ, 100µF, 70Ω coil). Used as reference denominator in Test 1 Result formula.
Col I — Measured Value | [MANUAL INPUT: numeric] | [5 empty] | Actual measured value from test instrument (multimeter, LCR meter, oscilloscope). Compared against Col H in the Test 1 Result formula.
Col M — Test 2 Result | [MANUAL INPUT: Pass / Fail] | [5 empty] | Result of functional test (e.g., logic function, switching response, communication test). Must be observed directly — cannot be calculated from a single measurement.
ℹ Test instruments used: Multimeter (Sanwa CD771), LCR Meter, Oscilloscope (GDS-1022), DC Power Supply (HY3005F), Function Generator. Tolerance threshold ±10% aligned with IEC 60068 incoming inspection practices.

## Elek_Digital
PRAKTIKUM ELEKTRONIKA DIGITAL — COMPONENT QC INSPECTION LOG
Lab Sensor & Sistem Telekontrol (SSTK) ¦ Dept. Teknik Nuklir & Teknik Fisika, UGM ¦ ~20 Mahasiswa/Sesi
Test Equipment: Multimeter (Sanwa CD771), Digital Logic Probe, DC Power Supply (HY3005F), Function Generator
No. | Part Number | Component Description | Batch / Lot No. | Inspection Date | Qty | Visual Inspection | Spec / Rated Value | Measured Value | Functional Test 1 | Test 1 Result | Functional Test 2 | Test 2 Result | Overall Result | Remarks / Disposition
⚠ Columns shaded yellow = manual input per inspection session ¦ Columns shaded gray = auto-calculated by formula ¦ Overall Result = auto (PASS only if ALL tests pass)
Inspection Session: Sesi A — Week 1 ¦ Batch: LOT-A-2024
1 | IC-74LS00 | NAND Gate (Quad) | LOT-A-949 | 05-Apr-2026 | 22 | Pass | 5 | 4.7 | VCC Test | Pass | Logic Function Test | Pass | PASS | Release to practicum
2 | IC-74LS08 | AND Gate (Quad) | LOT-A-934 | 05-Apr-2026 | 24 | Pass | 5 | 4.83 | VCC Test | Pass | Logic Function Test | Pass | PASS | Release to practicum
3 | IC-74LS86 | XOR Gate (Quad) | LOT-A-507 | 05-Apr-2026 | 21 | Pass | 5 | 4.78 | VCC Test | Pass | Logic Function Test | Pass | PASS | Release to practicum
4 | IC-74LS74 | D Flip-Flop | LOT-A-942 | 05-Apr-2026 | 25 | Pass | 5 | 4.83 | VCC Test | Pass | Toggle Test | Pass | PASS | Release to practicum
5 | IC-74LS163 | 4-bit Counter | LOT-A-686 | 06-Apr-2026 | 21 | Pass | 5 | 5.55 | VCC Test | Fail | Count Sequence Test | Pass | FAIL | Quarantine — do not issue to students
6 | IC-555 | Timer IC | LOT-A-579 | 06-Apr-2026 | 21 | Pass | 15 | 13.39 | VCC Test | Fail | Oscillation Test | Pass | FAIL | Quarantine — do not issue to students
7 | RLY-G2R-DC5 | Relay SPDT 5VDC | LOT-A-369 | 06-Apr-2026 | 20 | Pass | 70 | 63.86 | Coil R Meas | Pass | Switching Test | Pass | PASS | Release to practicum
8 | RLY-SRD-12V | Relay SPDT 12VDC | LOT-A-882 | 06-Apr-2026 | 22 | Pass | 400 | 373.14 | Coil R Meas | Pass | Switching Test | Pass | PASS | Release to practicum
9 | SW-TACT-6mm | Tactile Switch | LOT-A-353 | 07-Apr-2026 | 20 | Pass | 0.5 | 0.49 | Contact R | Pass | Continuity Test | Pass | PASS | Release to practicum
10 | SW-SLIDE | Slide Switch SPDT | LOT-A-591 | 07-Apr-2026 | 24 | Pass | 0.5 | 0.56 | Contact R | Fail | Continuity Test | Pass | FAIL | Quarantine — do not issue to students
11 | LED-RED-5mm | LED Red 5mm | LOT-A-811 | 07-Apr-2026 | 24 | Fail | 2 | 2.11 | Vf Meas | Fail | Brightness Check | Pass | FAIL | Quarantine — do not issue to students
12 | LED-GRN-5mm | LED Green 5mm | LOT-A-726 | 07-Apr-2026 | 19 | Pass | 2.1 | 2.35 | Vf Meas | Fail | Brightness Check | Pass | FAIL | Quarantine — do not issue to students
Inspection Session: Sesi B — Week 2 ¦ Batch: LOT-B-2025
1 | IC-74LS00 | NAND Gate (Quad) | LOT-B-876 | 12-Apr-2026 | 22 | Pass | 5 | 4.44 | VCC Test | Fail | Logic Function Test | Pass | FAIL | Quarantine — do not issue to students
2 | IC-74LS08 | AND Gate (Quad) | LOT-B-786 | 12-Apr-2026 | 24 | Pass | 5 | 4.63 | VCC Test | Pass | Logic Function Test | Pass | PASS | Release to practicum
3 | IC-74LS86 | XOR Gate (Quad) | LOT-B-271 | 12-Apr-2026 | 25 | Pass | 5 | 4.97 | VCC Test | Pass | Logic Function Test | Pass | PASS | Release to practicum
4 | IC-74LS74 | D Flip-Flop | LOT-B-227 | 12-Apr-2026 | 19 | Pass | 5 | 4.72 | VCC Test | Pass | Toggle Test | Pass | PASS | Release to practicum
5 | IC-74LS163 | 4-bit Counter | LOT-B-328 | 13-Apr-2026 | 20 | Pass | 5 | 4.68 | VCC Test | Pass | Count Sequence Test | Pass | PASS | Release to practicum
6 | IC-555 | Timer IC | LOT-B-936 | 13-Apr-2026 | 20 | Pass | 15 | 16.11 | VCC Test | Pass | Oscillation Test | Pass | PASS | Release to practicum
7 | RLY-G2R-DC5 | Relay SPDT 5VDC | LOT-B-446 | 13-Apr-2026 | 21 | Pass | 70 | 72.18 | Coil R Meas | Pass | Switching Test | Pass | PASS | Release to practicum
8 | RLY-SRD-12V | Relay SPDT 12VDC | LOT-B-120 | 13-Apr-2026 | 20 | Pass | 400 | 388.67 | Coil R Meas | Pass | Switching Test | Pass | PASS | Release to practicum
9 | SW-TACT-6mm | Tactile Switch | LOT-B-618 | 14-Apr-2026 | 20 | Pass | 0.5 | 0.47 | Contact R | Pass | Continuity Test | Pass | PASS | Release to practicum
10 | SW-SLIDE | Slide Switch SPDT | LOT-B-819 | 14-Apr-2026 | 20 | Pass | 0.5 | 0.51 | Contact R | Pass | Continuity Test | Pass | PASS | Release to practicum
11 | LED-RED-5mm | LED Red 5mm | LOT-B-791 | 14-Apr-2026 | 19 | Pass | 2 | 1.88 | Vf Meas | Pass | Brightness Check | Pass | PASS | Release to practicum
12 | LED-GRN-5mm | LED Green 5mm | LOT-B-533 | 14-Apr-2026 | 23 | Pass | 2.1 | 1.86 | Vf Meas | Fail | Brightness Check | Pass | FAIL | Quarantine — do not issue to students
Inspection Session: Sesi C — Week 3 ¦ Batch: LOT-C-2026
1 | IC-74LS00 | NAND Gate (Quad) | LOT-C-572 | 19-Apr-2026 | 25 | Pass | 5 | 4.98 | VCC Test | Pass | Logic Function Test | Pass | PASS | Release to practicum
2 | IC-74LS08 | AND Gate (Quad) | LOT-C-665 | 19-Apr-2026 | 24 | Pass | 5 | 5.22 | VCC Test | Pass | Logic Function Test | Pass | PASS | Release to practicum
3 | IC-74LS86 | XOR Gate (Quad) | LOT-C-766 | 19-Apr-2026 | 23 | Pass | 5 | 5.17 | VCC Test | Pass | Logic Function Test | Pass | PASS | Release to practicum
4 | IC-74LS74 | D Flip-Flop | LOT-C-261 | 19-Apr-2026 | 22 | Fail | 5 | 5.28 | VCC Test | Fail | Toggle Test | Pass | FAIL | Quarantine — do not issue to students
5 | IC-74LS163 | 4-bit Counter | LOT-C-301 | 20-Apr-2026 | 20 | Pass | 5 | 4.96 | VCC Test | Pass | Count Sequence Test | Pass | PASS | Release to practicum
6 | IC-555 | Timer IC | LOT-C-852 | 20-Apr-2026 | 19 | Pass | 15 | 13.38 | VCC Test | Fail | Oscillation Test | Pass | FAIL | Quarantine — do not issue to students
7 | RLY-G2R-DC5 | Relay SPDT 5VDC | LOT-C-825 | 20-Apr-2026 | 21 | Pass | 70 | 72.87 | Coil R Meas | Pass | Switching Test | Pass | PASS | Release to practicum
8 | RLY-SRD-12V | Relay SPDT 12VDC | LOT-C-639 | 20-Apr-2026 | 23 | Pass | 400 | 423.94 | Coil R Meas | Pass | Switching Test | Pass | PASS | Release to practicum
9 | SW-TACT-6mm | Tactile Switch | LOT-C-771 | 21-Apr-2026 | 18 | Pass | 0.5 | 0.48 | Contact R | Pass | Continuity Test | Pass | PASS | Release to practicum
10 | SW-SLIDE | Slide Switch SPDT | LOT-C-793 | 21-Apr-2026 | 21 | Pass | 0.5 | 0.53 | Contact R | Pass | Continuity Test | Pass | PASS | Release to practicum
11 | LED-RED-5mm | LED Red 5mm | LOT-C-416 | 21-Apr-2026 | 20 | Pass | 2 | 1.91 | Vf Meas | Pass | Brightness Check | Pass | PASS | Release to practicum
12 | LED-GRN-5mm | LED Green 5mm | LOT-C-737 | 21-Apr-2026 | 20 | Pass | 2.1 | 2.12 | Vf Meas | Pass | Brightness Check | Pass | PASS | Release to practicum
Inspection Session: Sesi D — Week 4 ¦ Batch: LOT-D-2027
1 | IC-74LS00 | NAND Gate (Quad) | LOT-D-154 | 26-Apr-2026 | 21 | Pass | 5 | 5.15 | VCC Test | Pass | Logic Function Test | Pass | PASS | Release to practicum
2 | IC-74LS08 | AND Gate (Quad) | LOT-D-291 | 26-Apr-2026 | 20 | Pass | 5 | 5.37 | VCC Test | Pass | Logic Function Test | Pass | PASS | Release to practicum
3 | IC-74LS86 | XOR Gate (Quad) | LOT-D-202 | 26-Apr-2026 | 24 | Pass | 5 | 5.17 | VCC Test | Pass | Logic Function Test | Pass | PASS | Release to practicum
4 | IC-74LS74 | D Flip-Flop | LOT-D-304 | 26-Apr-2026 | 23 | Pass | 5 | 4.56 | VCC Test | Pass | Toggle Test | Fail | FAIL | Quarantine — do not issue to students
5 | IC-74LS163 | 4-bit Counter | LOT-D-868 | 27-Apr-2026 | 25 | Pass | 5 | 4.7 | VCC Test | Pass | Count Sequence Test | Pass | PASS | Release to practicum
6 | IC-555 | Timer IC | LOT-D-435 | 27-Apr-2026 | 24 | Fail | 15 | 15.62 | VCC Test | Fail | Oscillation Test | Pass | FAIL | Quarantine — do not issue to students
7 | RLY-G2R-DC5 | Relay SPDT 5VDC | LOT-D-832 | 27-Apr-2026 | 19 | Pass | 70 | 69.03 | Coil R Meas | Pass | Switching Test | Pass | PASS | Release to practicum
8 | RLY-SRD-12V | Relay SPDT 12VDC | LOT-D-620 | 27-Apr-2026 | 24 | Pass | 400 | 441.9 | Coil R Meas | Fail | Switching Test | Pass | FAIL | Quarantine — do not issue to students
9 | SW-TACT-6mm | Tactile Switch | LOT-D-608 | 28-Apr-2026 | 25 | Pass | 0.5 | 0.47 | Contact R | Pass | Continuity Test | Fail | FAIL | Quarantine — do not issue to students
10 | SW-SLIDE | Slide Switch SPDT | LOT-D-960 | 28-Apr-2026 | 22 | Pass | 0.5 | 0.5 | Contact R | Pass | Continuity Test | Fail | FAIL | Quarantine — do not issue to students
11 | LED-RED-5mm | LED Red 5mm | LOT-D-926 | 28-Apr-2026 | 23 | Pass | 2 | 1.94 | Vf Meas | Pass | Brightness Check | Pass | PASS | Release to practicum
12 | LED-GRN-5mm | LED Green 5mm | LOT-D-171 | 28-Apr-2026 | 25 | Pass | 2.1 | 2.25 | Vf Meas | Pass | Brightness Check | Pass | PASS | Release to practicum
SESSION SUMMARY
Metric | [5 empty] | Value | Total Records | PASS | FAIL | Pass Rate
[7 empty] | 48 | 34 | 14 | 70.83%
📋 FORMULA & CALCULATION NOTES — ELEK DIGITAL
Column / Field | Calculation Logic
Col K — Test 1 Result | [Formula] IF(g="Fail","Fail", IF(ABS(i-h)/h ≤ 10%, "Pass", "Fail")),TRUE())),TRUE())),TRUE())),TRUE())) | [5 empty] | AUTO. Fails if Visual (G) already failed. Otherwise, passes only if measured value (I) is within ±10% of the rated spec (H). Tolerance 10% covers standard component variability.
Col N — Overall Result | [Formula] IF(AND(g="Pass", k="Pass", m="Pass"), "PASS", "FAIL") | [5 empty] | AUTO. Component passes only if ALL three tests pass: Visual (G), Test 1 (K), and Test 2 (M). Any single failure triggers FAIL — consistent with incoming inspection standards.
Col O — Remarks / Disposition | [Formula] IF(n="PASS", "Release to practicum", "Quarantine — do not issue") | [5 empty] | AUTO. Directly derived from Overall Result. PASS components are cleared for student use. FAIL components are quarantined and tracked in the Defect Log sheet.
Col G — Visual Inspection | [MANUAL INPUT: Pass / Fail] | [5 empty] | Inspector physically checks component body, pin integrity, marking, and housing. Recorded as Pass or Fail. Basis for all downstream formula calculations.
Col H — Spec / Rated Value | [MANUAL INPUT: numeric] | [5 empty] | Rated value from component datasheet or procurement spec (e.g., 10kΩ, 100µF, 70Ω coil). Used as reference denominator in Test 1 Result formula.
Col I — Measured Value | [MANUAL INPUT: numeric] | [5 empty] | Actual measured value from test instrument (multimeter, LCR meter, oscilloscope). Compared against Col H in the Test 1 Result formula.
Col M — Test 2 Result | [MANUAL INPUT: Pass / Fail] | [5 empty] | Result of functional test (e.g., logic function, switching response, communication test). Must be observed directly — cannot be calculated from a single measurement.
ℹ Test instruments used: Multimeter (Sanwa CD771), LCR Meter, Oscilloscope (GDS-1022), DC Power Supply (HY3005F), Function Generator. Tolerance threshold ±10% aligned with IEC 60068 incoming inspection practices.

## Sistem_Instr
PRAKTIKUM SISTEM INSTRUMENTASI — COMPONENT QC INSPECTION LOG
Lab Sensor & Sistem Telekontrol (SSTK) ¦ Dept. Teknik Nuklir & Teknik Fisika, UGM ¦ ~20 Mahasiswa/Sesi
Test Equipment: Multimeter (Sanwa CD771), DC Power Supply, Oscilloscope (GDS-1022), Arduino Uno (test harness), Serial Monitor
No. | Part Number | Component Description | Batch / Lot No. | Inspection Date | Qty | Visual Inspection | Spec / Rated Value | Measured Value | Functional Test 1 | Test 1 Result | Functional Test 2 | Test 2 Result | Overall Result | Remarks / Disposition
⚠ Columns shaded yellow = manual input per inspection session ¦ Columns shaded gray = auto-calculated by formula ¦ Overall Result = auto (PASS only if ALL tests pass)
Inspection Session: Sesi A — Week 1 ¦ Batch: LOT-A-2024
1 | SEN-LM35 | Temp Sensor LM35 | LOT-A-980 | 05-Apr-2026 | 22 | Pass | 10 | 10.7 | Output @25°C | Pass | Linearity Test | Pass | PASS | Release to practicum
2 | SEN-DS18B20 | 1-Wire Temp Sensor | LOT-A-758 | 05-Apr-2026 | 22 | Pass | 0.5 | 0.54 | Measured Err | Pass | Communication Test | Pass | PASS | Release to practicum
3 | SEN-BMP180 | Pressure Sensor | LOT-A-214 | 05-Apr-2026 | 24 | Fail | 1 | 1.09 | Measured Err | Fail | I2C Comm Test | Pass | FAIL | Quarantine — do not issue to students
4 | SEN-LDR | LDR Photo Sensor | LOT-A-267 | 05-Apr-2026 | 22 | Pass | 10 | 9.25 | Dark R Meas | Pass | Light Response Test | Pass | PASS | Release to practicum
5 | SEN-PIR-HC | PIR Motion Sensor | LOT-A-161 | 06-Apr-2026 | 22 | Pass | 5 | 4.6 | Trigger Test | Pass | Detection Range | Pass | PASS | Release to practicum
6 | SEN-US-HC04 | Ultrasonic HC-SR04 | LOT-A-898 | 06-Apr-2026 | 22 | Pass | 400 | 444.68 | Accuracy | Fail | Echo Response Test | Pass | FAIL | Quarantine — do not issue to students
7 | MOD-ESP32 | ESP32 MCU Module | LOT-A-937 | 06-Apr-2026 | 20 | Pass | 3.3 | 2.97 | Boot Test | Pass | GPIO Continuity | Pass | PASS | Release to practicum
8 | MOD-ADS1115 | ADC Module 16-bit | LOT-A-210 | 06-Apr-2026 | 24 | Pass | 3.3 | 3.53 | I2C Comm Test | Pass | Conversion Test | Pass | PASS | Release to practicum
9 | MOD-4-20mA | 4-20mA Transmitter | LOT-A-556 | 07-Apr-2026 | 20 | Pass | 20 | 18.77 | Zero @4mA | Pass | Span @20mA | Fail | FAIL | Quarantine — do not issue to students
10 | XDCR-K-TC | K-Type Thermocouple | LOT-A-881 | 07-Apr-2026 | 21 | Pass | 2.5 | 2.36 | Output @25°C | Pass | Insulation Test | Fail | FAIL | Quarantine — do not issue to students
11 | XDCR-PT100 | RTD PT100 | LOT-A-475 | 07-Apr-2026 | 19 | Pass | 100 | 91.49 | Measured R @RT | Pass | Continuity Test | Pass | PASS | Release to practicum
12 | MOD-OLED | OLED Display 0.96" | LOT-A-570 | 07-Apr-2026 | 19 | Pass | 3.3 | 3.48 | Boot Display Test | Pass | I2C Comm Test | Pass | PASS | Release to practicum
Inspection Session: Sesi B — Week 2 ¦ Batch: LOT-B-2025
1 | SEN-LM35 | Temp Sensor LM35 | LOT-B-940 | 12-Apr-2026 | 21 | Pass | 10 | 9.92 | Output @25°C | Pass | Linearity Test | Pass | PASS | Release to practicum
2 | SEN-DS18B20 | 1-Wire Temp Sensor | LOT-B-336 | 12-Apr-2026 | 19 | Pass | 0.5 | 0.49 | Measured Err | Pass | Communication Test | Pass | PASS | Release to practicum
3 | SEN-BMP180 | Pressure Sensor | LOT-B-888 | 12-Apr-2026 | 22 | Pass | 1 | 1.06 | Measured Err | Pass | I2C Comm Test | Pass | PASS | Release to practicum
4 | SEN-LDR | LDR Photo Sensor | LOT-B-943 | 12-Apr-2026 | 20 | Pass | 10 | 10.93 | Dark R Meas | Pass | Light Response Test | Pass | PASS | Release to practicum
5 | SEN-PIR-HC | PIR Motion Sensor | LOT-B-348 | 13-Apr-2026 | 20 | Pass | 5 | 4.64 | Trigger Test | Pass | Detection Range | Pass | PASS | Release to practicum
6 | SEN-US-HC04 | Ultrasonic HC-SR04 | LOT-B-382 | 13-Apr-2026 | 24 | Pass | 400 | 439.53 | Accuracy | Pass | Echo Response Test | Pass | PASS | Release to practicum
7 | MOD-ESP32 | ESP32 MCU Module | LOT-B-508 | 13-Apr-2026 | 19 | Pass | 3.3 | 3.03 | Boot Test | Pass | GPIO Continuity | Pass | PASS | Release to practicum
8 | MOD-ADS1115 | ADC Module 16-bit | LOT-B-845 | 13-Apr-2026 | 19 | Pass | 3.3 | 3.11 | I2C Comm Test | Pass | Conversion Test | Pass | PASS | Release to practicum
9 | MOD-4-20mA | 4-20mA Transmitter | LOT-B-746 | 14-Apr-2026 | 21 | Pass | 20 | 17.99 | Zero @4mA | Fail | Span @20mA | Pass | FAIL | Quarantine — do not issue to students
10 | XDCR-K-TC | K-Type Thermocouple | LOT-B-228 | 14-Apr-2026 | 24 | Pass | 2.5 | 2.68 | Output @25°C | Pass | Insulation Test | Pass | PASS | Release to practicum
11 | XDCR-PT100 | RTD PT100 | LOT-B-576 | 14-Apr-2026 | 25 | Pass | 100 | 101.54 | Measured R @RT | Pass | Continuity Test | Pass | PASS | Release to practicum
12 | MOD-OLED | OLED Display 0.96" | LOT-B-598 | 14-Apr-2026 | 21 | Pass | 3.3 | 3.36 | Boot Display Test | Pass | I2C Comm Test | Pass | PASS | Release to practicum
Inspection Session: Sesi C — Week 3 ¦ Batch: LOT-C-2026
1 | SEN-LM35 | Temp Sensor LM35 | LOT-C-960 | 19-Apr-2026 | 24 | Pass | 10 | 9.33 | Output @25°C | Pass | Linearity Test | Pass | PASS | Release to practicum
2 | SEN-DS18B20 | 1-Wire Temp Sensor | LOT-C-127 | 19-Apr-2026 | 22 | Pass | 0.5 | 0.53 | Measured Err | Pass | Communication Test | Pass | PASS | Release to practicum
3 | SEN-BMP180 | Pressure Sensor | LOT-C-861 | 19-Apr-2026 | 22 | Pass | 1 | 1.05 | Measured Err | Pass | I2C Comm Test | Pass | PASS | Release to practicum
4 | SEN-LDR | LDR Photo Sensor | LOT-C-336 | 19-Apr-2026 | 25 | Pass | 10 | 10.35 | Dark R Meas | Pass | Light Response Test | Pass | PASS | Release to practicum
5 | SEN-PIR-HC | PIR Motion Sensor | LOT-C-453 | 20-Apr-2026 | 22 | Pass | 5 | 5.24 | Trigger Test | Pass | Detection Range | Pass | PASS | Release to practicum
6 | SEN-US-HC04 | Ultrasonic HC-SR04 | LOT-C-445 | 20-Apr-2026 | 25 | Pass | 400 | 444.43 | Accuracy | Fail | Echo Response Test | Pass | FAIL | Quarantine — do not issue to students
7 | MOD-ESP32 | ESP32 MCU Module | LOT-C-308 | 20-Apr-2026 | 19 | Pass | 3.3 | 3.25 | Boot Test | Pass | GPIO Continuity | Pass | PASS | Release to practicum
8 | MOD-ADS1115 | ADC Module 16-bit | LOT-C-249 | 20-Apr-2026 | 23 | Pass | 3.3 | 3.35 | I2C Comm Test | Pass | Conversion Test | Pass | PASS | Release to practicum
9 | MOD-4-20mA | 4-20mA Transmitter | LOT-C-902 | 21-Apr-2026 | 18 | Pass | 20 | 18.51 | Zero @4mA | Pass | Span @20mA | Pass | PASS | Release to practicum
10 | XDCR-K-TC | K-Type Thermocouple | LOT-C-706 | 21-Apr-2026 | 20 | Pass | 2.5 | 2.6 | Output @25°C | Pass | Insulation Test | Pass | PASS | Release to practicum
11 | XDCR-PT100 | RTD PT100 | LOT-C-260 | 21-Apr-2026 | 23 | Pass | 100 | 97.57 | Measured R @RT | Pass | Continuity Test | Pass | PASS | Release to practicum
12 | MOD-OLED | OLED Display 0.96" | LOT-C-418 | 21-Apr-2026 | 23 | Pass | 3.3 | 3.48 | Boot Display Test | Pass | I2C Comm Test | Pass | PASS | Release to practicum
Inspection Session: Sesi D — Week 4 ¦ Batch: LOT-D-2027
1 | SEN-LM35 | Temp Sensor LM35 | LOT-D-603 | 26-Apr-2026 | 25 | Pass | 10 | 9.86 | Output @25°C | Pass | Linearity Test | Pass | PASS | Release to practicum
2 | SEN-DS18B20 | 1-Wire Temp Sensor | LOT-D-180 | 26-Apr-2026 | 22 | Pass | 0.5 | 0.53 | Measured Err | Pass | Communication Test | Fail | FAIL | Quarantine — do not issue to students
3 | SEN-BMP180 | Pressure Sensor | LOT-D-272 | 26-Apr-2026 | 19 | Pass | 1 | 1.08 | Measured Err | Pass | I2C Comm Test | Pass | PASS | Release to practicum
4 | SEN-LDR | LDR Photo Sensor | LOT-D-706 | 26-Apr-2026 | 20 | Pass | 10 | 9.3 | Dark R Meas | Pass | Light Response Test | Pass | PASS | Release to practicum
5 | SEN-PIR-HC | PIR Motion Sensor | LOT-D-973 | 27-Apr-2026 | 21 | Pass | 5 | 4.81 | Trigger Test | Pass | Detection Range | Pass | PASS | Release to practicum
6 | SEN-US-HC04 | Ultrasonic HC-SR04 | LOT-D-976 | 27-Apr-2026 | 21 | Pass | 400 | 409.19 | Accuracy | Pass | Echo Response Test | Pass | PASS | Release to practicum
7 | MOD-ESP32 | ESP32 MCU Module | LOT-D-588 | 27-Apr-2026 | 23 | Pass | 3.3 | 3.15 | Boot Test | Pass | GPIO Continuity | Pass | PASS | Release to practicum
8 | MOD-ADS1115 | ADC Module 16-bit | LOT-D-496 | 27-Apr-2026 | 25 | Pass | 3.3 | 3.13 | I2C Comm Test | Pass | Conversion Test | Fail | FAIL | Quarantine — do not issue to students
9 | MOD-4-20mA | 4-20mA Transmitter | LOT-D-479 | 28-Apr-2026 | 25 | Pass | 20 | 22.19 | Zero @4mA | Fail | Span @20mA | Pass | FAIL | Quarantine — do not issue to students
10 | XDCR-K-TC | K-Type Thermocouple | LOT-D-252 | 28-Apr-2026 | 20 | Pass | 2.5 | 2.65 | Output @25°C | Pass | Insulation Test | Pass | PASS | Release to practicum
11 | XDCR-PT100 | RTD PT100 | LOT-D-789 | 28-Apr-2026 | 21 | Pass | 100 | 93.38 | Measured R @RT | Pass | Continuity Test | Pass | PASS | Release to practicum
12 | MOD-OLED | OLED Display 0.96" | LOT-D-507 | 28-Apr-2026 | 24 | Pass | 3.3 | 3.18 | Boot Display Test | Pass | I2C Comm Test | Pass | PASS | Release to practicum
SESSION SUMMARY
Metric | [5 empty] | Value | Total Records | PASS | FAIL | Pass Rate
[7 empty] | 48 | 39 | 9 | 81.25%
📋 FORMULA & CALCULATION NOTES — SISTEM INSTR
Column / Field | Calculation Logic
Col K — Test 1 Result | [Formula] IF(g="Fail","Fail", IF(ABS(i-h)/h ≤ 10%, "Pass", "Fail")),TRUE())),TRUE())),TRUE())),TRUE())) | [5 empty] | AUTO. Fails if Visual (G) already failed. Otherwise, passes only if measured value (I) is within ±10% of the rated spec (H). Tolerance 10% covers standard component variability.
Col N — Overall Result | [Formula] IF(AND(g="Pass", k="Pass", m="Pass"), "PASS", "FAIL") | [5 empty] | AUTO. Component passes only if ALL three tests pass: Visual (G), Test 1 (K), and Test 2 (M). Any single failure triggers FAIL — consistent with incoming inspection standards.
Col O — Remarks / Disposition | [Formula] IF(n="PASS", "Release to practicum", "Quarantine — do not issue") | [5 empty] | AUTO. Directly derived from Overall Result. PASS components are cleared for student use. FAIL components are quarantined and tracked in the Defect Log sheet.
Col G — Visual Inspection | [MANUAL INPUT: Pass / Fail] | [5 empty] | Inspector physically checks component body, pin integrity, marking, and housing. Recorded as Pass or Fail. Basis for all downstream formula calculations.
Col H — Spec / Rated Value | [MANUAL INPUT: numeric] | [5 empty] | Rated value from component datasheet or procurement spec (e.g., 10kΩ, 100µF, 70Ω coil). Used as reference denominator in Test 1 Result formula.
Col I — Measured Value | [MANUAL INPUT: numeric] | [5 empty] | Actual measured value from test instrument (multimeter, LCR meter, oscilloscope). Compared against Col H in the Test 1 Result formula.
Col M — Test 2 Result | [MANUAL INPUT: Pass / Fail] | [5 empty] | Result of functional test (e.g., logic function, switching response, communication test). Must be observed directly — cannot be calculated from a single measurement.
ℹ Test instruments used: Multimeter (Sanwa CD771), LCR Meter, Oscilloscope (GDS-1022), DC Power Supply (HY3005F), Function Generator. Tolerance threshold ±10% aligned with IEC 60068 incoming inspection practices.

## Jar_Komunikasi
PRAKTIKUM JARINGAN KOMUNIKASI — COMPONENT QC INSPECTION LOG
Lab Sensor & Sistem Telekontrol (SSTK) ¦ Dept. Teknik Nuklir & Teknik Fisika, UGM ¦ ~20 Mahasiswa/Sesi
Test Equipment: Multimeter (Sanwa CD771), DC Power Supply, Spectrum Analyzer (basic), Serial Terminal (Arduino IDE), Oscilloscope
No. | Part Number | Component Description | Batch / Lot No. | Inspection Date | Qty | Visual Inspection | Spec / Rated Value | Measured Value | Functional Test 1 | Test 1 Result | Functional Test 2 | Test 2 Result | Overall Result | Remarks / Disposition
⚠ Columns shaded yellow = manual input per inspection session ¦ Columns shaded gray = auto-calculated by formula ¦ Overall Result = auto (PASS only if ALL tests pass)
Inspection Session: Sesi A — Week 1 ¦ Batch: LOT-A-2024
1 | MOD-NRF24L01 | RF Module 2.4GHz | LOT-A-785 | 05-Apr-2026 | 18 | Pass | 3.3 | 3.05 | Power Test | Pass | TX/RX Comm Test | Pass | PASS | Release to practicum
2 | MOD-HC-05 | Bluetooth Module | LOT-A-407 | 05-Apr-2026 | 21 | Pass | 5 | 5.24 | AT Cmd Response | Pass | Pairing Test | Pass | PASS | Release to practicum
3 | MOD-SIM800L | GSM/GPRS Module | LOT-A-608 | 05-Apr-2026 | 18 | Pass | 4 | 4.17 | AT Cmd Response | Pass | SIM Card Detection | Pass | PASS | Release to practicum
4 | MOD-LORA32 | LoRa Module 433MHz | LOT-A-761 | 05-Apr-2026 | 21 | Fail | 3.3 | 3.08 | TX Power | Fail | RX Sensitivity | Pass | FAIL | Quarantine — do not issue to students
5 | MOD-MAX485 | RS-485 Transceiver | LOT-A-742 | 06-Apr-2026 | 21 | Pass | 5 | 5.09 | Power Test | Pass | Tx/Rx Enable Test | Pass | PASS | Release to practicum
6 | MOD-MCP2515 | CAN Bus Controller | LOT-A-604 | 06-Apr-2026 | 20 | Pass | 5 | 4.98 | SPI Comm Test | Pass | CAN Frame Test | Pass | PASS | Release to practicum
7 | CON-RJ45 | RJ45 Connector Cat5e | LOT-A-748 | 06-Apr-2026 | 25 | Pass | 8 | 7.14 | Continuity | Fail | Crimp Quality | Pass | FAIL | Quarantine — do not issue to students
8 | CBL-USB-A | USB-A to USB-B 1m | LOT-A-763 | 06-Apr-2026 | 20 | Pass | 4 | 3.94 | Pin Continuity | Pass | Data Transfer Test | Fail | FAIL | Quarantine — do not issue to students
9 | REL-G5V-DC5 | Signal Relay 5VDC | LOT-A-421 | 07-Apr-2026 | 20 | Pass | 50 | 52.03 | Coil R Meas | Pass | Switching Test | Pass | PASS | Release to practicum
10 | ANT-2.4G | 2.4GHz PCB Antenna | LOT-A-938 | 07-Apr-2026 | 23 | Pass | 2 | 2.05 | Connector Check | Pass | Return Loss Test | Pass | PASS | Release to practicum
11 | TB-2P-5MM | Terminal Block 2P | LOT-A-609 | 07-Apr-2026 | 24 | Pass | 10 | 8.84 | Contact R | Fail | Screw Torque Test | Pass | FAIL | Quarantine — do not issue to students
12 | TB-3P-5MM | Terminal Block 3P | LOT-A-483 | 07-Apr-2026 | 23 | Pass | 10 | 9.72 | Contact R | Pass | Screw Torque Test | Pass | PASS | Release to practicum
Inspection Session: Sesi B — Week 2 ¦ Batch: LOT-B-2025
1 | MOD-NRF24L01 | RF Module 2.4GHz | LOT-B-714 | 12-Apr-2026 | 21 | Pass | 3.3 | 3.69 | Power Test | Fail | TX/RX Comm Test | Pass | FAIL | Quarantine — do not issue to students
2 | MOD-HC-05 | Bluetooth Module | LOT-B-774 | 12-Apr-2026 | 18 | Pass | 5 | 5.27 | AT Cmd Response | Pass | Pairing Test | Pass | PASS | Release to practicum
3 | MOD-SIM800L | GSM/GPRS Module | LOT-B-592 | 12-Apr-2026 | 18 | Pass | 4 | 3.92 | AT Cmd Response | Pass | SIM Card Detection | Pass | PASS | Release to practicum
4 | MOD-LORA32 | LoRa Module 433MHz | LOT-B-385 | 12-Apr-2026 | 18 | Fail | 3.3 | 3.62 | TX Power | Fail | RX Sensitivity | Pass | FAIL | Quarantine — do not issue to students
5 | MOD-MAX485 | RS-485 Transceiver | LOT-B-993 | 13-Apr-2026 | 25 | Pass | 5 | 4.63 | Power Test | Pass | Tx/Rx Enable Test | Pass | PASS | Release to practicum
6 | MOD-MCP2515 | CAN Bus Controller | LOT-B-285 | 13-Apr-2026 | 19 | Pass | 5 | 5 | SPI Comm Test | Pass | CAN Frame Test | Fail | FAIL | Quarantine — do not issue to students
7 | CON-RJ45 | RJ45 Connector Cat5e | LOT-B-200 | 13-Apr-2026 | 22 | Fail | 8 | 7.98 | Continuity | Fail | Crimp Quality | Fail | FAIL | Quarantine — do not issue to students
8 | CBL-USB-A | USB-A to USB-B 1m | LOT-B-892 | 13-Apr-2026 | 18 | Pass | 4 | 3.88 | Pin Continuity | Pass | Data Transfer Test | Pass | PASS | Release to practicum
9 | REL-G5V-DC5 | Signal Relay 5VDC | LOT-B-883 | 14-Apr-2026 | 21 | Pass | 50 | 55.94 | Coil R Meas | Fail | Switching Test | Pass | FAIL | Quarantine — do not issue to students
10 | ANT-2.4G | 2.4GHz PCB Antenna | LOT-B-684 | 14-Apr-2026 | 23 | Pass | 2 | 1.96 | Connector Check | Pass | Return Loss Test | Fail | FAIL | Quarantine — do not issue to students
11 | TB-2P-5MM | Terminal Block 2P | LOT-B-185 | 14-Apr-2026 | 19 | Pass | 10 | 9.82 | Contact R | Pass | Screw Torque Test | Pass | PASS | Release to practicum
12 | TB-3P-5MM | Terminal Block 3P | LOT-B-776 | 14-Apr-2026 | 24 | Pass | 10 | 10.8 | Contact R | Pass | Screw Torque Test | Pass | PASS | Release to practicum
Inspection Session: Sesi C — Week 3 ¦ Batch: LOT-C-2026
1 | MOD-NRF24L01 | RF Module 2.4GHz | LOT-C-616 | 19-Apr-2026 | 18 | Pass | 3.3 | 3.51 | Power Test | Pass | TX/RX Comm Test | Pass | PASS | Release to practicum
2 | MOD-HC-05 | Bluetooth Module | LOT-C-591 | 19-Apr-2026 | 20 | Pass | 5 | 4.96 | AT Cmd Response | Pass | Pairing Test | Pass | PASS | Release to practicum
3 | MOD-SIM800L | GSM/GPRS Module | LOT-C-293 | 19-Apr-2026 | 20 | Pass | 4 | 3.69 | AT Cmd Response | Pass | SIM Card Detection | Pass | PASS | Release to practicum
4 | MOD-LORA32 | LoRa Module 433MHz | LOT-C-306 | 19-Apr-2026 | 22 | Pass | 3.3 | 2.94 | TX Power | Fail | RX Sensitivity | Pass | FAIL | Quarantine — do not issue to students
5 | MOD-MAX485 | RS-485 Transceiver | LOT-C-503 | 20-Apr-2026 | 18 | Pass | 5 | 5.45 | Power Test | Pass | Tx/Rx Enable Test | Pass | PASS | Release to practicum
6 | MOD-MCP2515 | CAN Bus Controller | LOT-C-677 | 20-Apr-2026 | 18 | Pass | 5 | 5.4 | SPI Comm Test | Pass | CAN Frame Test | Fail | FAIL | Quarantine — do not issue to students
7 | CON-RJ45 | RJ45 Connector Cat5e | LOT-C-393 | 20-Apr-2026 | 21 | Pass | 8 | 7.66 | Continuity | Pass | Crimp Quality | Pass | PASS | Release to practicum
8 | CBL-USB-A | USB-A to USB-B 1m | LOT-C-548 | 20-Apr-2026 | 21 | Fail | 4 | 3.81 | Pin Continuity | Fail | Data Transfer Test | Pass | FAIL | Quarantine — do not issue to students
9 | REL-G5V-DC5 | Signal Relay 5VDC | LOT-C-729 | 21-Apr-2026 | 21 | Pass | 50 | 46.6 | Coil R Meas | Pass | Switching Test | Pass | PASS | Release to practicum
10 | ANT-2.4G | 2.4GHz PCB Antenna | LOT-C-982 | 21-Apr-2026 | 18 | Pass | 2 | 1.99 | Connector Check | Pass | Return Loss Test | Pass | PASS | Release to practicum
11 | TB-2P-5MM | Terminal Block 2P | LOT-C-206 | 21-Apr-2026 | 24 | Pass | 10 | 10.5 | Contact R | Pass | Screw Torque Test | Pass | PASS | Release to practicum
12 | TB-3P-5MM | Terminal Block 3P | LOT-C-347 | 21-Apr-2026 | 24 | Pass | 10 | 9.83 | Contact R | Pass | Screw Torque Test | Pass | PASS | Release to practicum
Inspection Session: Sesi D — Week 4 ¦ Batch: LOT-D-2027
1 | MOD-NRF24L01 | RF Module 2.4GHz | LOT-D-363 | 26-Apr-2026 | 20 | Fail | 3.3 | 3.45 | Power Test | Fail | TX/RX Comm Test | Fail | FAIL | Quarantine — do not issue to students
2 | MOD-HC-05 | Bluetooth Module | LOT-D-710 | 26-Apr-2026 | 19 | Pass | 5 | 4.96 | AT Cmd Response | Pass | Pairing Test | Pass | PASS | Release to practicum
3 | MOD-SIM800L | GSM/GPRS Module | LOT-D-611 | 26-Apr-2026 | 18 | Pass | 4 | 3.71 | AT Cmd Response | Pass | SIM Card Detection | Pass | PASS | Release to practicum
4 | MOD-LORA32 | LoRa Module 433MHz | LOT-D-996 | 26-Apr-2026 | 23 | Pass | 3.3 | 3.49 | TX Power | Pass | RX Sensitivity | Pass | PASS | Release to practicum
5 | MOD-MAX485 | RS-485 Transceiver | LOT-D-821 | 27-Apr-2026 | 22 | Pass | 5 | 5.57 | Power Test | Fail | Tx/Rx Enable Test | Pass | FAIL | Quarantine — do not issue to students
6 | MOD-MCP2515 | CAN Bus Controller | LOT-D-510 | 27-Apr-2026 | 19 | Pass | 5 | 5.57 | SPI Comm Test | Fail | CAN Frame Test | Pass | FAIL | Quarantine — do not issue to students
7 | CON-RJ45 | RJ45 Connector Cat5e | LOT-D-876 | 27-Apr-2026 | 23 | Pass | 8 | 7.94 | Continuity | Pass | Crimp Quality | Pass | PASS | Release to practicum
8 | CBL-USB-A | USB-A to USB-B 1m | LOT-D-858 | 27-Apr-2026 | 25 | Pass | 4 | 3.61 | Pin Continuity | Pass | Data Transfer Test | Pass | PASS | Release to practicum
9 | REL-G5V-DC5 | Signal Relay 5VDC | LOT-D-453 | 28-Apr-2026 | 23 | Pass | 50 | 45.03 | Coil R Meas | Pass | Switching Test | Pass | PASS | Release to practicum
10 | ANT-2.4G | 2.4GHz PCB Antenna | LOT-D-591 | 28-Apr-2026 | 18 | Pass | 2 | 1.77 | Connector Check | Fail | Return Loss Test | Pass | FAIL | Quarantine — do not issue to students
11 | TB-2P-5MM | Terminal Block 2P | LOT-D-456 | 28-Apr-2026 | 21 | Pass | 10 | 10.38 | Contact R | Pass | Screw Torque Test | Pass | PASS | Release to practicum
12 | TB-3P-5MM | Terminal Block 3P | LOT-D-205 | 28-Apr-2026 | 19 | Fail | 10 | 8.81 | Contact R | Fail | Screw Torque Test | Fail | FAIL | Quarantine — do not issue to students
SESSION SUMMARY
Metric | [5 empty] | Value | Total Records | PASS | FAIL | Pass Rate
[7 empty] | 48 | 30 | 18 | 62.5%
📋 FORMULA & CALCULATION NOTES — JAR KOMUNIKASI
Column / Field | Calculation Logic
Col K — Test 1 Result | [Formula] IF(g="Fail","Fail", IF(ABS(i-h)/h ≤ 10%, "Pass", "Fail")),TRUE())),TRUE())),TRUE())),TRUE())) | [5 empty] | AUTO. Fails if Visual (G) already failed. Otherwise, passes only if measured value (I) is within ±10% of the rated spec (H). Tolerance 10% covers standard component variability.
Col N — Overall Result | [Formula] IF(AND(g="Pass", k="Pass", m="Pass"), "PASS", "FAIL") | [5 empty] | AUTO. Component passes only if ALL three tests pass: Visual (G), Test 1 (K), and Test 2 (M). Any single failure triggers FAIL — consistent with incoming inspection standards.
Col O — Remarks / Disposition | [Formula] IF(n="PASS", "Release to practicum", "Quarantine — do not issue") | [5 empty] | AUTO. Directly derived from Overall Result. PASS components are cleared for student use. FAIL components are quarantined and tracked in the Defect Log sheet.
Col G — Visual Inspection | [MANUAL INPUT: Pass / Fail] | [5 empty] | Inspector physically checks component body, pin integrity, marking, and housing. Recorded as Pass or Fail. Basis for all downstream formula calculations.
Col H — Spec / Rated Value | [MANUAL INPUT: numeric] | [5 empty] | Rated value from component datasheet or procurement spec (e.g., 10kΩ, 100µF, 70Ω coil). Used as reference denominator in Test 1 Result formula.
Col I — Measured Value | [MANUAL INPUT: numeric] | [5 empty] | Actual measured value from test instrument (multimeter, LCR meter, oscilloscope). Compared against Col H in the Test 1 Result formula.
Col M — Test 2 Result | [MANUAL INPUT: Pass / Fail] | [5 empty] | Result of functional test (e.g., logic function, switching response, communication test). Must be observed directly — cannot be calculated from a single measurement.
ℹ Test instruments used: Multimeter (Sanwa CD771), LCR Meter, Oscilloscope (GDS-1022), DC Power Supply (HY3005F), Function Generator. Tolerance threshold ±10% aligned with IEC 60068 incoming inspection practices.

## Defect Log
COMPONENT DEFECT LOG — ALL PRACTIKUM
All FAIL records consolidated here. Input manually when Overall Result = FAIL. Disposition, root cause, and corrective action tracked per defect.
No. | Date | Part Number | Component | Practikum | Defect Type | Failed Test | Root Cause Analysis | Corrective Action | PIC | Status | Disposition
1 | 22-Apr-25 | R-10kΩ-CFR | Resistor 10kΩ | Elek Analog | Measured value OOSpec | Tolerance Test | Tolerance drift >5% — possible aged batch | Return to procurement; replace with new batch | Lab Staff | Closed | Quarantine — returned
2 | 23-Apr-25 | IC-74LS00 | NAND Gate IC | Elek Digital | Logic function failure | Logic Function Test | IC damaged — likely ESD during storage | Update anti-static storage SOP; reorder replacement | Lab Staff | Closed | Quarantine — disposed
3 | 24-Apr-25 | SEN-LM35 | Temp Sensor LM35 | Sistem Instr | Output deviation >10% | Output @25°C Test | Bent VCC pin — physical damage pre-inspection | Inspect all LM35 units for pin integrity before issuing | Lab Staff | Closed | Quarantine — disposed
4 | 25-Apr-25 | MOD-HC-05 | Bluetooth Module | Jar Komunikasi | No AT command response | AT Cmd Response | Firmware corrupted — likely improper power-down | Re-flash firmware; add power-down SOP to module handling | Lab Staff | In Progress | Hold — pending re-flash
5 | 28-Apr-25 | RLY-G2R-DC5 | Relay SPDT 5V | Elek Digital | Coil resistance OOSpec | Switching Test | Manufacturing variation — coil R 78Ω vs 70Ω spec | Issue CCAR to supplier; update incoming spec tolerance | Lab Staff | Open | Quarantine — return to vendor
6 | 29-Apr-25 | C-100µF-EC | Capacitor 100µF | Elek Analog | ESR too high | ESR Test | Electrolyte dry-out — old batch past shelf life | Check batch date on all capacitors; replace if >2yr old | Lab Staff | Closed | Quarantine — disposed
7 | 30-Apr-25 | SEN-HC-SR04 | Ultrasonic Sensor | Sistem Instr | Echo response failure | Echo Response Test | Transducer membrane damaged — physical impact | Improve storage packaging; add visual inspection step | Lab Staff | Closed | Quarantine — disposed
8 | 01-May-25 | SW-SLIDE | Slide Switch SPDT | Elek Digital | Contact resistance high | Continuity Test | Oxidized contacts — storage humidity exposure | Store switches in sealed anti-humidity bags | Lab Staff | Open | Quarantine — pending cleaning
DEFECT SUMMARY |  |  |  | Total Defects | Closed | In Progress | Open | Closure Rate
|  |  |  | 8 | 5 | 1 | 2 | 0%

## Defect Analysis & Trend
DEFECT ANALYSIS — PARETO & MONTHLY TREND ¦ Lab SSTK UGM
All defect counts are formula-driven from inspection log sheets ¦ Monthly trend: input columns B,C,D,E,F,I only — rest auto-calculated
| Pareto: ELEKTRONIKA ANALOG | [5 empty] | Pareto: ELEKTRONIKA DIGITAL
| Defect Category | Count | % of Total | Cumulative % | Signal |  | Defect Category | Count | % of Total | Cumulative % | Signal
| Measured Value OOSpec | 3 | 50% | 50% | ▲ High |  | Measured Value OOSpec | 8 | 57.14% | 57.14% | ▲ High
| Visual Defect | 0 | 0% | 50% | — |  | Visual Defect | 3 | 21.43% | 78.57% | ▲ High
| Functional Test 2 Fail | 3 | 50% | 100% | ▲ High |  | Functional Test 2 Fail | 3 | 21.43% | 100% | ▲ High
| Both Tests Fail | 0 | 0% | 100% | — |  | Both Tests Fail | 0 | 0% | 100% | —
| TOTAL | 6 |  |  |  |  | TOTAL | 14
| Pareto: SISTEM INSTRUMENTASI | [5 empty] | Pareto: JARINGAN KOMUNIKASI
| Defect Category | Count | % of Total | Cumulative % | Signal |  | Defect Category | Count | % of Total | Cumulative % | Signal
| Sensor Output OOSpec | 4 | 44.44% | 44.44% | ▲ High |  | Module Comm Failure | 8 | 38.1% | 38.1% | ▲ High
| Visual Defect | 1 | 11.11% | 55.56% | ● Normal |  | Visual / Physical Defect | 6 | 28.57% | 66.67% | ▲ High
| Comm Test Fail | 4 | 44.44% | 100% | ▲ High |  | Secondary Test Fail | 4 | 19.05% | 85.71% | ▲ High
| Both Tests Fail | 0 | 0% | 100% | — |  | Both Tests Fail | 3 | 14.29% | 100% | ▲ High
| TOTAL | 9 |  |  |  |  | TOTAL | 21
Monthly Defect Rate Trend — Lab SSTK (All Practikum Combined)
Input B,C,D,E,F,I only. G=Total Defects(=D+E+F), H=Defect Rate(=G/C), J=Status(auto)
| Month [INPUT] | Total Inspected [INPUT] | Electronic Defects [INPUT] | Sensor/Comm Defects [INPUT] | Visual Defects [INPUT] | Total Defects [=D+E+F] | Defect Rate [=G/C] | Target Rate [INPUT] | Status [AUTO]
| Jan-25 | 320 | 14 | 8 | 6 | 28 | 8.75% | 0.05% | Above Target
| Feb-25 | 320 | 13 | 7 | 5 | 25 | 7.812% | 0.05% | Above Target
| Mar-25 | 336 | 11 | 6 | 4 | 21 | 6.25% | 0.05% | Above Target
| Apr-25 | 336 | 10 | 6 | 3 | 19 | 5.655% | 0.05% | Above Target
| May-25 | 352 | 8 | 5 | 3 | 16 | 4.545% | 0.05% | Above Target
| Jun-25 | 352 | 7 | 4 | 3 | 14 | 3.977% | 0.05% | Above Target
| Jul-25 | 368 | 6 | 4 | 2 | 12 | 3.261% | 0.05% | Above Target
| Aug-25 | 368 | 5 | 3 | 2 | 10 | 2.717% | 0.05% | Above Target
| Sep-25 | 368 | 5 | 3 | 2 | 10 | 2.717% | 0.05% | Above Target
| Oct-25 | 384 | 4 | 3 | 1 | 8 | 2.083% | 0.05% | Above Target
| Nov-25 | 384 | 4 | 2 | 1 | 7 | 1.823% | 0.05% | Above Target
| Dec-25 | 400 | 3 | 2 | 1 | 6 | 1.5% | 0.05% | Above Target
| TOTAL/AVG | 4288 | 90 | 53 | 33 | 176 | 4.104%
📋 FORMULA & CALCULATION NOTES — PARETO DEFECT CATEGORIES
Column / Field | Calculation Logic
Count (Pareto table) | [Formula] COUNTIFS(sheet!g:g,"Pass", sheet!k:k,"Fail") | [7 empty] | AUTO. Counts rows where Visual passed but Test 1 failed — isolating measurement OOSpec defects. Each defect category uses a different COUNTIFS combination across the inspection log sheet.
% of Total | 0 | [7 empty] | AUTO. Defect count for this category divided by total defects for that practikum. Expressed as percentage. Shows relative weight of each defect type — Pareto principle.
Cumulative % | [Formula] previous_row_cumulative + this_row_% | [7 empty] | AUTO. Running sum of percentages from top to bottom. When cumulative % reaches 80%, the categories above it are the 'vital few' — prioritised for corrective action per Pareto analysis.
Signal Column | [Formula] IF(count=0,"—", IF(count>2,"▲ High","● Normal")) | [7 empty] | AUTO. Quick visual flag. ▲ High = more than 2 defects of this type (requires investigation). ● Normal = 1-2 defects (monitor). — = no defect recorded.
ℹ Pareto counts reference live inspection data in the 4 practikum sheets. Adding new inspection records automatically updates all pareto counts.
📋 FORMULA & CALCULATION NOTES — MONTHLY TREND CALCULATIONS
Column / Field | Calculation Logic
Col G — Total Defects | [Formula] d + e + f | [7 empty] | AUTO. Sum of Electronic Defects (D) + Sensor/Comm Defects (E) + Visual Defects (F). Three input categories chosen to align with lab's 4 practikum domains.
Col H — Defect Rate | [Formula] g / c (total defects / total inspected) | [7 empty] | AUTO. Overall defect rate for the month. Expressed as percentage. Key KPI tracked against target. If H > I (target), Status turns red automatically.
Col J — Status | [Formula] IF(h ≤ I, "On Target", "Above Target"),TRUE()),TRUE()),TRUE()),TRUE()) | [7 empty] | AUTO. Green 'On Target' if defect rate meets or beats the monthly target. Red 'Above Target' triggers review of that month's inspection records.
Cols B,C,D,E,F,I | [MANUAL INPUT per month] | [7 empty] | Month label, total components inspected, defect counts by category, and target rate are entered manually — typically compiled from the 4 inspection log sheets at month-end.
ℹ Monthly trend data is compiled manually from inspection logs at end of each month. Target rate 5.0% aligned with lab quality improvement programme benchmark.